## Edit history
Version | date | ID | Changes | Remark
V0.1 | 2025.05.19 |  | First version V0.1 | Liu Yuzhi
V0.2 | June 12, 2025 |  | Add electronic seal event records to the [Records] section.
| June 18, 2025 | B070H | Alarm status words now include meter removal and strong magnetic interference status bits.
| June 24, 2025 | 70C1H | Delete data ID 70C1H and configure calendar type (Gregorian calendar or Persian calendar).
| June 27, 2025 | 70DAH | The pump control switch data ID has been changed from 4800H to 70DAH.
| 2025.06.30 | 70DEH | Add a new one-click production cleanup command (backup) for clearing production parameters on the main control board.
| 2025.07.11 | 70BFH | Remove 70BFH daylight saving time enable
| 2025.07.14 | 1060H | Add command to read remote communication status
| 2025.07.15 | 70C0H | 1. The description for the Daylight Saving Time parameter configuration in version 1.70C0H has been added: Unit: 10 minutes, Range: [-12, 12]. Setting the time to 0 indicates that Daylight Saving Time is not activated. 2. "Remaining Allowed Usage" has been changed to "Remaining Allowed Usage for This Period". 3. The read/write attribute of the remaining amount (remaining recharge amount) has been changed from RW to R.
| 2025.07.25 | B070H | Delete status bits: Water meter (flow meter) low battery event alarm (lithium battery voltage), pressure acquisition abnormality, pressure exceeding upper limit, pressure exceeding lower limit. Add status bits: Relay fault status.
| 2025.08.05 | 0018H | Delete this data ID; 2G modules do not support reading module model.
| 2025.08.05 | 70C5H | The number of bytes for time zone settings has been changed from 1 to 2, and the unit has been changed from 10 minutes to 1 minute. The range is now [-720, 840].
| 2025.08.08 |  | Delete terminal identification sheet
| 2025.08.11 | B06BH | Change the water usage log data ID in minutes (restore to original). Delete 70C6H and 70C7H.
|  | 70C0H | Daylight saving time parameter configuration data ID is placed in a composite data object sheet.
| August 12, 2025 | 70EEH | Add the IDs of the 10 most recently occurring events.
|  | 210BH | For 2G communication, add 210BH CSQ-GPRS, delete 0033H RSRP, 0034H SINR. All data IDs containing sub-items 0033H and 0034H need to be re-applied for: delete A324H, A325H, and A326H, and re-applied for communication event data IDs 70EFH, 70F0H, and 70F1H; delete 70C3H and re-applied for data ID 70F2H.
|  | 0039H | The length has been changed to 4 bytes, and the unit has been changed to L/h.
|  | 70F4H 70F5H | Changes to data IDs for water usage cycles and allowed water consumption within a cycle: Delete 70BCH, add 70F4H data ID; Delete 70BDH, add 70F5H data ID.
|  | 4705H 0049H | Delete Data ID
| 2025.08.15 | B038H | Modify the value range and modify the default value.
|  | B047H | Modify the value range
|  | B048H | Modify default value
|  | B04FH | Modify default value
|  | B040H | New
|  | B041H | New
|  | B050H | Delete ID
|  | B051H | Delete ID
V0.3 | 2025.08.26 | 70F5H | Remove "Default 0xFFFFFFFFFF is unlimited"
|  | 2012H | The default customer APN is empty (all 0s).
|  | 0061H | New
| 2025.9.23 | 21B4H | Remove the "1 - Exit without modification" function.
|  | B04DH | Change read/write attribute to read-only
V0.4 | 2026.1.29 | 7129H | New | Ding Zhongwa

## Simple Data Objects
Notes: 1. For data ID read/write permissions in different modes and scenarios, if √, the permission is consistent with the data ID itself; if R, it indicates read-only permission; if W, it indicates write-only permission; if blank, it indicates no read/write permission; if empty in all scenarios, it means that the data ID does not support standalone read/write operations; 2. Unless otherwise specified, big-endian format is used. | [7 empty] | Data terminal |  |  | Analysis | Remark
Command Function | inheritID | Chinese name | Data types | Magnification | Read and write attributes | length | describe | Production mode | Customer Model
[8 empty] | local | local | remote
[8 empty] | Business dataUnencrypted | Business dataencryption | Business dataencryption
Measurement and settlement | B023H | Positive cumulative amount | Unsigned integer | ×1 | RW | 8 | Unit L, unsigned integer | √ | √ |  | Measurement | 1) The cumulative amount for both directions is reset to zero at the factory. 2) Parameters written in customer mode require the electronic seal to be removed.
Measurement and settlement | B024H | Reverse cumulative | Unsigned integer | ×1 | RW | 8 | Unit L, unsigned integer | √ | √
Terminal operating parameters | 0002H | clock | BCD code integer | ×1 | RW | 6 |  | √ | √ | √ | business | The first frame contains
Terminal operating parameters | 001BH | Event Code | Unsigned integer | ×1 | R | 2 | Unsigned integer, used only as a sub-item ID in a record, does not support individual access.
Terminal operating parameters | 21B0H | Event Code Subdivision Event | Unsigned integer | ×1 | R | 1 | An unsigned integer, used only as an ID for a sub-item in a record, and does not support individual access.
Terminal operating parameters | 70B5H | Pump operating time | Unsigned integer | ×1 | RW | 4 | Unit h | √ | R | R |  | The first frame contains
Terminal configuration parameters | 4211H | Main controller version number | BCD code integer | ×1 | R | 4 | BCD code | √ | √ | √ | business
Terminal configuration parameters | 2001H | Software version (metric version) | BCD code integer | ×1 | R | 4 | BCD code | √ | √ |  | Measurement | Not open to customers
Terminal configuration parameters | A100H | Work mode | Unsigned integer | ×1 | RW | 1 | 0: Production mode; 1: Customer mode | √ | √ |  | business | Cut-only mode, no need to pass through to the metering board.
power supply | 0004H | Battery voltage (mains power) | Unsigned integer | ×1000 | R | 2 | Unsigned integer (HEX), unit 0.001V | √ | √ | √ | business | The first frame contains
power supply | 0006H | Backup power voltage | Unsigned integer | ×1000 | R | 2 | Unsigned integer (HEX), unit 0.001V | √ | √ | √ | business | The first frame contains
Measurement and settlement | B03FH | Instantaneous flow | Signed integers | ×100 | R | 8 | Signed integer, unit: 0.01 L/h, for example, 1600, is 16.00 L/h. | √ | √ |  | Measurement
Production | 2197H | One-click cleanup command for production | Signed integers | ×1 | W | 1 | Unsigned integer (HEX). Cleans up statistics in production mode, including event logs, water usage logs, runtime, etc. 1 command is valid. | √ |  |  | Measurement
Production | 70DEH | One-click cleanup command for production (backup) | Signed integers | ×1 | W | 1 | Unsigned integer (HEX). Cleans up statistics in production mode, including event logs, water usage logs, runtime, etc. 1 command is valid. | √ |  |  | business | Clearing freeze, events, remaining allowable usage for this period, additional usage for this period, remaining quantity, pump operating time, etc., for in-plant use only.
Terminal configuration parameters | B01BH | RS485 baud rate | Signed integers | ×1 | RW | 2 | The default value is 9600, in bps, with a range of 1200/2400/4800/9600.Synchronize with infrared communication baud rate settings | √ | √ |  | business
Terminal configuration parameters | 0066H | Local communication key (formerly known as: Maintain Default Key) | ASCII strings | ×1 | RW | 16 | 16 bytes | √ |  |  | business
Terminal configuration parameters | 0067H | Remote communication key (formerly known as: client default key) | ASCII strings | ×1 | RW | 16 | 16 bytes | √ | R | √ | business | The remote key is readable locally; the first 4 bytes and the last 4 bytes are hidden, while the middle 8 bytes are hidden.
Terminal operating parameters | 0039H | Average instantaneous flow(Average daily traffic) | Signed integers | ×1 | R | 4 | Average daily flow rate, in L/h. For example, 1600 represents 16.00 L/h. 0x7FFFFFFF indicates an anomaly. | √ | √ | √ | business | The first frame contains
Terminal configuration parameters | 70B6H | Valve/pump action configuration when the event occurs | Unsigned integer | ×1 | RW | 2 | 2-byte configuration bits, each bit represents: 0 - pump not shut off, 1 - pump off. Bit 0: Isolation door open, default 0 - pump not shut off. Bit 1: Exceeding maximum allowable water consumption, default 1 - pump off. Bit 2: Disconnecting the water meter power cable, default 1 - pump off. Bit 3: Disconnecting the pump control cable, default 1 - pump off (unused, no cable signal detected). Bit 4: Insufficient balance, default 1 - pump off. Bit 5: Meter disconnection, default 1 - pump off. Bit 6: Strong magnetic interference, default 1 - pump off. | √ | √ | √ | business
Terminal operating parameters | 70B9H | Remaining allowable usage for this period | Unsigned integer | ×1 | R | 5 | Unit 10L | √ | √ | √ |  | The first frame contains
Terminal operating parameters | 70BAH | Additional usage this period | Unsigned integer | ×1 | R | 5 | Unit 10L | √ | √ | √ |  | The first frame contains
Terminal operating parameters | 70BBH | Maximum daily traffic | Signed integers | ×1 | R | 3 | The unit is 10L/h, for example: 16 means 160L/h. | √ | √ | √ |  | The first frame contains
Terminal configuration parameters | 70BEH | Production date | BCD code integer | ×1 | RW | 3 | Production date: 0 if not set (year, month, day). | √ | R | R
Terminal configuration parameters | 70C5H | Time zone settings | Signed integers | ×1 | RW | 2 | The unit is min; for example, setting it to 60 means +60 min. The range is [-720, 840]. | √
Terminal configuration parameters | 70C1H | Configure the calendar type (Gregorian calendar or Persian calendar). | Unsigned integer | ×1 | RW | 1 | 0 - Gregorian calendar, 1 - Persian calendar (default is 1 - Persian calendar) | √ | √ | √
Terminal configuration parameters | 70DAH | Valve/Pump Control | Unsigned integer | ×1 | RW | 1 | 0 - Exit lock, 1 - Lock open, 2 - Lock closed, other values are invalid. | √ | √ | √
Terminal configuration parameters | 0027H | Remaining amount (remaining recharge amount) | Signed integers | ×1 | R | 4 | Unit L | R | R | R |  | The first frame contains
Terminal configuration parameters | 70C9H | STS Key Settings | HEX | ×1 | RW | 8 | The default form number is preceded by four digits, which are manually entered, e.g., 0000. Form number: 282007300167; STS key = 0000282007300167 | √
communication | 0028H | Start remote communication (start remote communication from local communication) | Unsigned integer | ×1 | W | 1 | 1 - Triggers a remote communication request; otherwise, it is invalid. This command is only valid for local communication (meaning local communication initiates remote communication). It is automatically cleared to 0 after this remote communication request is triggered. | √ |  |  | business
communication | 1060H | Read remote communication status | Unsigned integer | ×1 | R | 1 | 0: Not started; 1: Communicating; 2: Communication successful; 3: Communication failed;For internal factory use, data can be reported to the designated server; no data interaction is required. | √
Electronic Seal | 2079H | Electronic seal status | Unsigned integer | ×1 | R | 1 | Unsigned integer: 0: Seal not open; 1: Seal closed; 2: Status set. |  | √ |  | business | Not shown to customers
Electronic Seal | 21B2H | Unlock Hardware Seal Password | BCD code integer | ×1 | W | 2 | BCD code |  | √
Electronic Seal | 21B4H | Confirm or exit the sealing command | Unsigned integer | ×1 | W | 1 | 0: Confirm changes and exit.1: Exit without making any changes.Other reservations |  | √
Electronic Seal | 001DH | Modify Person Information | ASCII strings | ×1 | R | 10 | An ASCII string used to record information about the person who modified the electronic seal during the process. It is only used as a sub-item ID in the record and does not support individual access.
Electronic Seal | 21B5H | Electronic Seal Password | BCD code integer | ×1 | RW | 2 | The electronic seal password is only used as the ID for a sub-item in the record and will not be accessed independently.
communication | 210BH | CSQ-GPRS | Unsigned integer | ×1 | R | 1 | CSQ-GPRS | √ | √ | √ |  | The first frame contains
communication | 200EH | Peak shifting interval time (peak shifting coefficient) | Unsigned integer | ×1 | RW | 2 | The default value is 6.Data range 0 ~ 10, no peak shift at time 0.; | √ | √ | √ | business
communication | 0016H | IMEI (Module IMEI Number) | ASCII strings | ×1 | R | 15 | Module IMEI number | √ | √ | √ | business
communication | 0018H | Module model | ASCII strings | ×1 | R | 10 |  | √ | √ | √ | business | Not open to customers
communication | A013H | IMSI | ASCII strings | ×1 | R | 15 |  | √ | √ | √ | business
communication | 200AH | SIM card information (ICCID) | ASCII strings | ×1 | R | 20 | 20 bytes, corresponding to the SIM card's ICCID | √ | √ | √ | business
communication | 2012H | Customer APN | ASCII strings | ×1 | RW | 32 | If the value is less than 0, add 0s to the end; otherwise, leave the string empty (all zeros). | √ | √ | √ | business
communication | 0061H | Producing APN | ASCII strings | ×1 | RW | 50 | If the value is less than 0, add zeros to the end; the default value is "cmnet". | √ | √ | √ | business
Terminal configuration parameters | B055H | Freeze interval (minutes) | Unsigned integer | ×1 | RW | 1 | Values are limited to 30, 60, 90, and 120, in minutes, with a default of 60 minutes. | √ | √ | √ | business
Terminal operating parameters | B065H | Minute-by-minute freeze on positive cumulative volume (remote reporting) | Unsigned integer | ×1 | R | 5 | Unit: 10L, unsigned integer, 0xFFFFFFFFFF indicates an exception. As a sub-item ID, it will not be accessed individually. |  |  |  | business
Terminal operating parameters | B056H | Minute-by-minute freeze on reverse cumulative volume (remote reporting) | Unsigned integer | ×1 | R | 5 | Unit: 10L, unsigned integer, 0xFFFFFFFFFF indicates an exception. As a sub-item ID, it will not be accessed individually.
Terminal operating parameters | B077H | Daily frozen positive cumulative amount (remote reporting) | Unsigned integer | ×1 | R | 5 | Unit: 10L, unsigned integer, 0xFFFFFFFFFF indicates an exception. As a sub-item ID, it will not be accessed individually. |  |  |  |  | The first frame contains
Terminal operating parameters | B078H | Daily frozen reverse cumulative volume (remote reporting) | Unsigned integer | ×1 | R | 5 | Unit: 10L, unsigned integer, 0xFFFFFFFFFF indicates an exception. As a sub-item ID, it will not be accessed individually. |  |  |  |  | The first frame contains
Terminal operating parameters | B079H | Current positive cumulative amount (remotely reported) | Unsigned integer | ×1 | R | 5 | Unit: 10L, unsigned integer, 0xFFFFFFFFFF indicates an exception. As a sub-item ID, it will not be accessed individually. |  |  |  |  | The first frame contains
Terminal operating parameters | B07AH | Current reverse cumulative amount (remotely reported) | Unsigned integer | ×1 | R | 5 | Unit: 10L, unsigned integer, 0xFFFFFFFFFF indicates an exception. As a sub-item ID, it will not be accessed individually. |  |  |  |  | The first frame contains
Measurement and settlement | B067H | Instantaneous flow rate (remote reporting) | Signed integers | ×1 | R | 3 | The unit is 10L/h, for example: 16 represents 160L/h. This is used as a sub-item ID. |  |  |  | business | The first frame contains
Terminal configuration parameters | B064H | Terminal communication identification code | Unsigned integer | ×1 | R | 6 | Retrieve the last 6 bytes of the meter number |  |  |  | business | The first frame contains
Terminal configuration parameters | B038H | Pipe diameter | Unsigned integer | ×1 | RW | 2 | Unsigned integers (HEX)Default 1000Unit: 0.1mm. Diameter range: 250, 320, 400, 500, 650, 800, 1000, 1250, 1500, 2000, 2500, 3000, 3200, 3500, 4000, 4500, 5000, 6000, 7000, 8000, 9000, 10000 | √ | √ |  | Measurement | Writing parameters in client mode requires removing the electronic seal.
Terminal configuration parameters | A204H | Standard mode or test mode | Unsigned integer | ×1 | RW | 1 | Unsigned integer (HEX). 0 - Standard 1 - Test | √ | √ |  | Measurement
Terminal configuration parameters | 2192H | Water meter types (metering principle) | Unsigned integer | ×1 | R | 2 | Unsigned Integer (HEX) 1-Single-Jet Rotary Type 2-Multi-Jet Rotary Type 3-Ultrasonic 4-Electromagnetic 5-Insertion Ultrasonic Flow Meter | √ |  |  | Measurement
Dry/Wet type (referring to mechanical water meters): 1 - indicates dry type; 2 - indicates wet type; 0 - indicates this field is meaningless and will not be parsed.
Production | 21B7H | Start Zero Drift Self-Learning | Unsigned integer | ×1 | W | 1 | Unsigned integer (HEX). 0: Off 1: On | √
Production | 0029H | Water meter zero drift self-learning state | Unsigned integer | ×1 | R | 1 | 0 - Zero-drift self-learning not enabled or self-learning failed; 1 - Zero-drift self-learning in progress; 2 - Self-learning completed (successful). | √
Production | 0037H | Start of calibration form | Unsigned integer | ×1 | W | 1 | 1- The meter starts counting the average instantaneous flow; this command is invalid for other purposes. This command is only valid for local communication. Upon receiving this command, the meter resets the average instantaneous flow to zero. | √ |  |  | Measurement
Production | 0038H | End of Checklist | Unsigned integer | ×1 | W | 1 | 1 - This command is only valid for local communication and will not be used for other purposes. It ends the statistical calculation of average instantaneous traffic. | √
Production | B045H | Set the calibration flow direction | Unsigned integer | ×1 | RW | 1 | 0 indicates corrected flow, 1 indicates corrected reverse flow, other values are invalid. Default is 0. | √ |  |  | Measurement
Terminal configuration parameters | B019H | Wired pulse configuration | uint8_t | ×1 | RW | 1 | 0 indicates no configuration, 1 indicates configuration, and others are invalid. The default is 0. | √
Terminal configuration parameters | B01AH | Electronic watch pulse equivalent calibration | uint32_t | ×1 | RW | 4 | The default value is 100, the unit is ml, and the value range is 1-9999999. | √
Terminal configuration parameters | B047H | Overcurrent alarm threshold | Unsigned integer | ×1 | RW | 4 | The default value for electromagnetic water meters is 1.25*Q3, and the default value for electromagnetic flow meters is 1.25qmax. The configuration range is 1000 ~ 999999999, unit: L/h. | √
Terminal configuration parameters | 4704H | Overcurrent detection enable switch | Unsigned integer | ×1 | RW | 1 | Function enable switch: 0 - disabled; 1 - enabled. Default: 0. | √
Terminal configuration parameters | 4707H | Overcurrent alarm duration | Unsigned integer | ×1 | RW | 1 | Value range: 1~255, default parameter 30, unit: minutes; | √
Terminal configuration parameters | 2190H | Reflux detection time interval | Unsigned integer | ×1 | RW | 1 | Value range: 1~255, default parameter 3, unit: minutes; | √
Terminal configuration parameters | 202FH | Reverse current alarm enable | Unsigned integer | ×1 | RW | 1 | 0: Off; 1: On; Default is 0. | √
Terminal configuration parameters | B048H | Backflow alarm threshold | Unsigned integer | ×1 | RW | 4 | The default value for electromagnetic water meters is 0.1*Q3; the default value for electromagnetic flow meters is 0.1*qmax, in L/h. | √
Terminal configuration parameters | B049H | Reverse metering settings | uint8_t | ×1 | RW | 1 | 0 indicates not enabled, 1 indicates enabled, other values are invalid. Default is 1. | √
Terminal configuration parameters | B04AH | Flow direction settings | uint8_t | ×1 | RW | 1 | 0: Forward; 1: Reverse. Default is 0. | √
Terminal configuration parameters | B04BH | Air traffic control alarm threshold | uint16_t | ×1 | RW | 2 | Value range: 1-65535, default value1700300Unit mv | √
Terminal configuration parameters | B04CH | Excitation alarm threshold | uint16_t | ×1 | RW | 2 | Value range: 1-65535, default value: 50, unit: mv | √
Terminal configuration parameters | B04FH | Range ratio (Q3/Q1) | Unsigned integer | ×1 | RW | 2 | The value range is 1 to 65535.The default value is 25.(Electromagnetic water meter) | √
Terminal configuration parameters | B052H | Measurement and collection cycle | uint16_t | ×1 | RW | 2 | Value range: 1000~15000 (in increments of 1000), default 10000, unit: milliseconds. | √
Production | B07CH | Electronic watch calibration pulse output enable | Unsigned integer | ×1 | RW | 1 | Function enable switch: 0 - disabled; 1 - enabled, default is 0. | √
Terminal configuration parameters | B07FH | Cumulative start-up filter flow threshold | unsigned integers | ×100 | RW | 4 | Default value Qs, unit 0.01L/h | √
Terminal configuration parameters | 704CH | Dynamic zero point value | Signed integers | ×10000 | RW | 2 | Value range: -0.010m/s to 0.010m/s; default value: 0, unit: 0.0001m/s | √
Terminal configuration parameters | 704DH | Converter coefficients | Unsigned integer | ×10000 | RW | 2 | Default value is 10000, unit is 0.0001 | √
Terminal configuration parameters | 704EH | Full scale setting value | Unsigned integer | ×1 | RW | 4 | The default value for electromagnetic water meters is Q3, and the default value for electromagnetic flow meters is qmax, with units of 1L/h. | √
Terminal configuration parameters | 704FH | Empty pipe inspection value | Unsigned integer | ×1 | R | 2 | Initial value 0, unit mv | √
Production | 7050H | Electromagnetic water meter flow meter settings | Unsigned integer | ×1 | RW | 1 | 0: Electromagnetic water meter; 1: Electromagnetic flow meter. Default value: 0 | √

## Composite data objects
Notes: 1. For data ID read/write permissions in different modes and scenarios, if √, the permission is consistent with the data ID itself; if R, it indicates read-only permission; if W, it indicates write-only permission; if blank, it indicates no read/write permission. If all scenarios are blank, it means that the data ID does not support standalone read/write operations. 2. Composite data objects do not contain simple data objects. | [9 empty] | Data terminal |  |  | Analysis | Remark
Command Function | ID | Chinese name | Read and write attributes | describe | Sub-item Chinese name | Sub-item length | Data types | Magnification | Sub-item description | Production mode | Customer Model
[10 empty] | local | local | remote
[10 empty] | Business data is not encrypted | Business dataencryption | Business dataencryption
Terminal configuration parameters | 2039H | Composite form number(Used for handshakes) | R |  | Terminal type | 2 | Unsigned integer | ×1 | During the local communication session handshake, the system fills this field with 0, and the meter does not parse this field. | √ |  |  | business
[5 empty] | meter number length | 1 | Unsigned integer | ×1 | meter number length: If the highest bit is 1, it indicates that the meter number format is ASCII string, and the meter number length is 1~17; if the highest bit is 0, it indicates BCD, and the meter number length is 1~34; the default is BCD.
[5 empty] | meter No. | 17 | BCD code integer | ×1 | meter number, BCD code, maximum 34-bit BCD.
Production | A102H | Meter number parameters (production)(Used for the whole machine serial number) | RW |  |  | 1 | Unsigned integer | ×1 | meter number length: If the highest bit is 1, it indicates that the meter number format is ASCII string, and the meter number length is 1~17; if the highest bit is 0, it indicates BCD, and the meter number length is 1~34; the default is BCD.The meter number is 7 bytes: 00000000 (serial number) 09 (customer code) 01 (electromagnetic water meter) 03 (using the Iranian calendar year, this year is Iranian year 1403). The water meter serial number is 4 bytes (00000000-99999999). | √ | R | R | business
[6 empty] | 17 | BCD code integer | ×1 | The meter number content defaults to BCD, 34 bits, padded with zeros.
communication | B061H | Scheduled parameter upload (high-frequency scheduled upload) | RW |  | Scheduled upload interval | 1 | unsigned integers | ×1 | Unsigned integer (HEX), default parameter is 72, unit: 10 minutes. Data range: 1, 2, 3, 6, 12, 18, 24, 36, 48, 72, 144 (i.e., set to 10 minutes, 20 minutes, 30 minutes, 1 hour, 2 hours, 3 hours, 4 hours, 6 hours, 8 hours, 12 hours, 24 hours) | √ | √ | √ | business
[5 empty] | Upload start time | 2 | BCD | ×1 | BCD. hhmm indicates the start time of the scheduled upload. Default is 00:00.
communication | 2007H | Client server IPv4 address and port | RW |  | Collection service address IPv4 | 16 | ASCII code | ×1 | Data collection service address: padded with zeros if less than 16 bytes | √ | √ | √ | business
[5 empty] | Data collection server port | 2 | Unsigned integer | ×1 | Data collection server port
communication | 0065H | Production server IPv4 address and port | RW |  | Production service address IPv4 | 16 | ASCII code | ×1 | Production service address: padded with 0s if less than 16 bytes | √ |  |  | business
[5 empty] | Production server port | 2 | Unsigned integer | ×1 | Production server port
Terminal configuration parameters | 70C0H | Daylight Saving Time Parameter Configuration | RW |  | Daylight Saving Time | 6 | BCD code integer | ×1 | 3-byte BCD code, start month, day, and hour of Daylight Saving Time: MMDDHH 3-byte BCD code, end month, day, and hour of Daylight Saving Time: MMDDHH Default start time: March 22, 00:00 - September 21, 00:00 | √ | √ | √ | business
[5 empty] | Daylight Saving Time Adjustment | 1 | Signed integers | ×1 | 1-byte HEX, unit: 10 minutes, range [-12, 12]. For example: 0x0a represents +100 minutes, 0xFF represents -10 minutes. A time adjustment of 0 indicates that daylight saving time is not activated.
Terminal configuration parameters | 70F4H | Period start/end time | R |  | Period start/end time | 6 | BCD code integer | ×1 | Cycle start: MMDDHH; Cycle end: MMDDHH | √ | √ | √ | business
[5 empty] | Period start/end time | 6 | BCD code integer | ×1 | Cycle start: MMDDHH; Cycle end: MMDDHH
[5 empty] | Period start/end time | 6 | BCD code integer | ×1 | Cycle start: MMDDHH; Cycle end: MMDDHH
[5 empty] | Period start/end time | 6 | BCD code integer | ×1 | Cycle start: MMDDHH; Cycle end: MMDDHH
Terminal configuration parameters | 70F5H | Permissible water consumption within the cycle | R |  | Permissible water consumption within the cycle | 5 | Unsigned integer | ×1 | A unit of 10L can be used to set a usage limit for one cycle, for example, 270F, which is 99990L. | √ | √ | √ | business
[5 empty] | Permissible water consumption within the cycle | 5 | Unsigned integer | ×1 | A unit of 10L can be used to set a usage limit for one cycle, for example, 270F, which is 99990L.
[5 empty] | Permissible water consumption within the cycle | 5 | Unsigned integer | ×1 | A unit of 10L can be used to set a usage limit for one cycle, for example, 270F, which is 99990L.
[5 empty] | Permissible water consumption within the cycle | 5 | Unsigned integer | ×1 | A unit of 10L can be used to set a usage limit for one cycle, for example, 270F, which is 99990L.
Electronic Seal | 21B3H | Unseal the electronic seal | W |  | Operation mark | 1 | Unsigned integer | ×1 | 0: Internal commands, 1: External commands (meter-side commands can be ignored for now) |  | √ |  | business | Not reflected to customers
[5 empty] | password | 2 | BCD code integer | ×1 | Electronic Seal Password
[5 empty] | Modified by | 10 | ASCII strings | ×1 | Modifier information: 10-byte ASCII code
Terminal configuration parameters | B053H | Zero flow threshold | RW | The upper limit threshold must be greater than or equal to the lower limit threshold. |  | 4 | Unsigned integer magnification factor | ×100 | Unsigned integer (HEX), threshold value in test mode, unit 0.01L/h. Default value for electromagnetic water meters: 1/10 Q1; default value for electromagnetic flow meters: 0.03%qmax. | √ |  |  | Measurement
[6 empty] | 4 |  |  | Unsigned integer (HEX), threshold value in test mode, unit 0.01L/h. Default value for electromagnetic water meters: 1/10 Q1; default value for electromagnetic flow meters: 0.03%qmax.
[6 empty] | 4 |  |  | Unsigned integer (HEX), threshold value in standard mode, unit 0.01 L/h. Default value for electromagnetic water meters: 1/2 Q1; default value for electromagnetic flow meters: 0.15%qmax.
[6 empty] | 4 |  |  | Unsigned integer (HEX), threshold value in standard mode, unit 0.01L/h. Electromagnetic water meter default value 1/3 Q1; electromagnetic flow meter default value 0.15%qmax.
Production | B040H | Set forward flow calibration parameters | RW |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 0/uint32_t/*unit is 1.0E-7m/s*/ | √ | √ |  | Measurement | Writing parameters in client mode requires removing the electronic seal.
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 0 / uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 1/uint32_t /*unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 1/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 2/uint32_t /*Unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 2/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 3/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 3/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 4/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 4/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 5/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 5/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 6/uint32_t/*unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 6/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 7/uint32_t/*unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 7/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 8/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 8/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 9/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 9/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 10/uint32_t /*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 10/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 11/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 11/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 12/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 12/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 13/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 13/uint16_t /*Unit is 0.0001*/
|  | Set forward flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 14/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 14/uint16_t /*Unit is 0.0001*/
Production | B041H | Set reverse flow calibration parameters | RW |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 0/uint32_t/*unit is 1.0E-7m/s*/ | √ | √ |  | Measurement | Writing parameters in client mode requires removing the electronic seal.
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 0 / uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 1/uint32_t /*unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 1/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 2/uint32_t /*Unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 2/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 3/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 3/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 4/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 4/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 5/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 5/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 6/uint32_t/*unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 6/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | The original flow velocity is 7/uint32_t/*unit is 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 7/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 8/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 8/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 9/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 9/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 10/uint32_t /*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 10/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 11/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 11/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 12/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 12/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 13/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 13/uint16_t /*Unit is 0.0001*/
|  | Set reverse flow calibration parameters |  |  |  | 4 | Unsigned integer magnification factor | ×10000000 | Original flow velocity 14/uint32_t/*Unit: 1.0E-7m/s*/
[6 empty] | 2 | Unsigned integer magnification factor | ×10000 | Correction value 14/uint16_t /*Unit is 0.0001*/
Production | B04DH | Electromagnetic zero drift self-learning value | R |  | Zero drift value in standard mode | 4 | Signed integers | ×1 | The default value is 0. | √ |  |  | Measurement
[5 empty] | Zero drift value in test mode | 4 | Signed integers | ×1 | The default value is 0.
Production | 7061H | Forward and reverse air tube voltage values | R/W | Voltage values for forward and reverse empty tube detection | Forward empty tube voltage value | 2 | Unsigned integer | ×1 | The high 2 bytes store the positive open tube voltage value. | R |  |  | Measurement
[5 empty] | Reverse tube voltage value | 2 | Unsigned integer | ×1 | The lower 2 bytes store the reverse open tube voltage value.
Production | 7129H | upper and lower thresholds of tube voltage | RW | The upper threshold must be greater than or equal to the lower threshold. | upper limit threshold of tube voltage | 2 | Unsigned integer | ×1 | Value range 1-65535, default 1800, unit mV | √ |  |  | Measurement
[5 empty] | Lower threshold of tube voltage | 2 | Unsigned integer | ×1 | Value range 1-65535, default 700, unit mV

## Composite bit object
Note: 1. For data ID read/write permissions in different modes and scenarios, if √, the permission is consistent with the data ID itself; if R, it indicates read-only permission; if W, it indicates write-only permission; if blank, it indicates no read/write permission. If all scenarios are blank, it means that the data ID does not support standalone read/write operations. | [8 empty] | Data terminal
ID | English name | name | describe | byte sequence number | Bit width | Sub-item English name | Sub-item Chinese name | Sub-item description | Production mode | Customer Model
[9 empty] | local | local | remote
[9 empty] | Business data is not encrypted | Business dataencryption | Business dataencryption
B070H |  | Alarm status words | Flag: 1 - Valid; 0 - Invalid | 0 | bit0 |  | Storage error |  | √ | √ | √
[5 empty] | bit1 |  | Traffic collection anomaly
[5 empty] | bit2 |  | Reserved
[5 empty] | bit3 |  | Low battery event alarm (lithium battery voltage)
[5 empty] | bit4 |  | Power lockout event alarm (lithium battery voltage)
[5 empty] | bit5 |  | Reserved
[5 empty] | bit6 |  | Reserved
[5 empty] | bit7 |  | External power access sign
|  |  |  | 1 | bit0 |  | Water meter (flow meter) empty pipe alarm
[5 empty] | bit1 |  | Water meter (flow meter) excitation alarm
[5 empty] | bit2 |  | Low battery alarm for water meter (flow meter) (lithium battery voltage)
[5 empty] | bit3 |  | Low signal strength of water meter (flow meter)
[5 empty] | bit4 |  | Water meter (flow meter) measurement error
[5 empty] | bit5 |  | Water meter (flow meter) backflow alarm
[5 empty] | bit6 |  | Water meter (flow meter) alarm for abnormally high flow rate
[5 empty] | bit7 |  | Reserved
|  |  |  | 2 | bit0 |  | Pump switch status: 0 - Normally on, 1 - Off
[5 empty] | bit1 |  | Disassemble
[5 empty] | bit2 |  | strong magnetic interference
[5 empty] | bit3 |  | Relay fault status
[5 empty] | bit4 |  | Reserved
[5 empty] | bit5 |  | Reserved
[5 empty] | bit6 |  | Reserved
[5 empty] | bit7 |  | Reserved

## array of objects
Note: The object array will not be accessed individually, so there is no need to test the read and write operations of the object array. |  |  |  |  | Data terminal
ID | name | describe | Element ID | Number of elements | Production mode | Customer Model
[5 empty] | local | local | remote
[5 empty] | Business data is not encrypted | Business dataencryption | Business dataencryption
70C6H | Minute-by-minute water usage log positive cumulative amount | Includes minute-by-minute water usage logs for a specified day, with varying lengths and 12-48 data entries. | B023H | /
70C7H | Minute water usage log reverse cumulative amount | Includes minute-by-minute water usage logs for a specified day, with varying lengths and 12-48 data entries. | B024H | /
B02DH | Positive cumulative amount of daily water usage log | Includes the positive cumulative amount frozen daily on the 31st of the specified month. | B023H | 31
B02EH | Daily water usage log reverse cumulative amount | Includes the reverse cumulative amount frozen daily on the 31st of the specified month. | B024H | 31
B02FH | Positive cumulative monthly water usage log | Includes the positive cumulative amount frozen for each of the 12 months of the specified year. | B023H | 12
B030H | Monthly water usage log reverse cumulative amount | Includes monthly frozen reverse accumulation for the specified year (12 months). | B024H | 12

## Firmware upgrade
Note: 1. For data ID read and write permissions in different modes and scenarios, if it is √, the permission is consistent with the data ID itself; if it is R, it means read-only permission; if it is W, it means write-only permission; if it is blank, it means no read or write permission; if it is blank in all scenarios, it means that the data ID does not support standalone read and write operations. | [9 empty] | Data terminal
Command Function | ID | name | Read and write attributes | describe | Sub-item Chinese name | Sub-item length | Data types | Magnification | Sub-item description | Production mode | Customer Model
[10 empty] | local | local | remote
[10 empty] | Business data is not encrypted | Business dataencryption | Business dataencryption
Firmware upgrade | 4304 | Meter upgrade status | R |  |  | 4 | Unsigned integer | ×1 | The first byte indicates the firmware download status: 0 No firmware download; 1 Confirmed upgrade status; 2 Resuming interrupted download; 3 Firmware download complete, awaiting installation; 4 Firmware download failed; 5 Firmware installation successful; 6 Firmware installation failed. The second byte indicates a firmware download failure flag, which is only meaningful when the first byte is 4: 0 No firmware download or firmware download successful; 1 Firmware download failed, connection to upgrade server failed; 2 Firmware download failed, data interaction failed (other failures besides 1, 3, and 4); 3 Firmware download failed, firmware verification error; 4 Firmware download failed, firmware loading to Flash failed. Other bytes are reserved. |  | √ | √
Firmware upgrade | 4305 | Firmware information | RW | The server sends firmware upgrade information to the meter. |  | 4 | BCD code integer | ×1 | Firmware version number |  | √ | √
[6 empty] | 4 | Unsigned integer | ×1 | Firmware size
[6 empty] | 4 | Unsigned integer | ×1 | Firmware CRC32 verification
Firmware upgrade | 4306 | Firmware segmentation parameters | R | The instrument requests a firmware package from the server. |  | 4 | Unsigned integer | ×1 | Firmware package offset: refers to the offset relative to the entire firmware to be downloaded. For example, 0x400 indicates the 1Kth firmware file. |  | √ | √
[6 empty] | 4 | Unsigned integer | ×1 | Firmware package effective length
Firmware upgrade | 4307 | Firmware package data structure | W | The server responded with firmware package data. |  | 4 | HEX string | ×1 | Firmware package offset address: the offset address relative to the program. |  | √ | √
[6 empty] | 4 | HEX string | ×1 | Firmware package size
[6 empty] | N | HEX string | ×1 | Firmware data
[6 empty] | 2 | Unsigned integer | ×1 | Firmware package CRC16 check
Firmware upgrade | 430C | Server upgrade request | W | Server request for upgrade |  | 1 | Unsigned integer | ×1 | The upgrade server sends an update to table 0x00: Do not upgrade; 0x01 confirms the upgrade flag. |  | √ | √
[6 empty] | 4 | BCD code integer | ×1 | Target firmware version number
[6 empty] | 4 | BCD code integer | ×1 | Current firmware version number
Firmware upgrade | 430E | Upgrade status upload response | W | The server responded with ID 4304. |  | 1 | Unsigned integer | ×1 | 0x01 indicates that the upgrade server has confirmed receiving the upgrade status from the table. |  | √ | √
Firmware upgrade | B03CH | Force stop the upgrade | W | The server was forced to stop upgrading. |  | 1 | Unsigned integer | ×1 | 0x01This indicates that the server has been forced to stop the upgrade. |  | √ | √

## Record
Data terminal
Function Description | ID | Chinese name | describe | Time stamp and sub-item ID | Sub-item description | Production mode | Customer Model
[6 empty] | local | local | remote
[6 empty] | Business data is not encrypted | Business data encryption | Business data encryption
Communication network abnormal event log | 21A2H | Communication network abnormal event log |  | YYMMDDhhmmss |  | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
System error log -- Reset log | 100BH | System error log -- Reset log |  | YYMMDDhhmmss |  | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 001CH | Task ID
Alarm Event Log -- Abnormal Alarm Event Log | 21A3H | Alarm Event Log -- Abnormal Alarm Event Log |  | YYMMDDhhmmss |  | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
Business event log | 005AH | Business event log |  | YYMMDDhhmmss |  | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
Scheduled Remote Communication Log (GPRS) -- Scheduled Remote Communication Failure | 70EFH |  |  | 0002H | YYMMDDhhmmss | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
|  |  |  | 210BH | CSQ-GPRS
Manual remote communication log (GPRS) | 70F0H |  |  | 0002H | YYMMDDhhmmss | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
|  |  |  | 210BH | CSQ-GPRS
Scheduled remote transmission of communication logs not started (GPRS) | 70F1H |  |  | 0002H | YYMMDDhhmmss | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
|  |  |  | 210BH | CSQ-GPRS
Electronic seal modification of forward flow correction coefficient record | B042H | Electronic seal modification correction coefficient record | Length is 198 | YYMMDDhhmmss |  | √ | √ |  | Measurement
|  |  |  | B040H | Original correction coefficient
|  |  |  | B040H | Modified correction coefficient
|  |  |  | 21B5H | Unseal the electronic seal password (the high two bytes of the issued password).
|  |  |  | 001DH | Modify person's identity
Electronic seal modification reverse flow correction coefficient record | B043H | Electronic seal modification correction coefficient record | Length is 198 | YYMMDDhhmmss |  | √ | √ |  | Measurement
|  |  |  | B041H | Original correction coefficient
|  |  |  | B041H | Modified correction coefficient
|  |  |  | 21B5H | Unseal the electronic seal password (the high two bytes of the issued password).
|  |  |  | 001DH | Modify person's identity
Electronic seal modification positive cumulative record | B039H | Electronic seal modification positive cumulative record | 34 | YYMMDDhhmmss |  | √ | √ |  | Measurement
|  |  |  | B023H | Positive cumulative amount
|  |  |  | B023H | Modified positive cumulative amount
|  |  |  | 21B5H | Unseal the electronic seal password (the high two bytes of the issued password).
|  |  |  | 001DH | Modify person's identity
Electronic seal modification reverse cumulative record | B03AH | Electronic seal modification reverse cumulative record | 34 | YYMMDDhhmmss |  | √ | √ |  | Measurement
|  |  |  | B024H | Reverse cumulative
|  |  |  | B024H | Modified reverse cumulative
|  |  |  | 21B5H | Unseal the electronic seal password (the high two bytes of the issued password).
|  |  |  | 001DH | Modify person's identity
Electronic seal modification of pipe diameter record | B03BH | Electronic seal modification of pipe diameter record | 22 | YYMMDDhhmmss |  | √ | √ |  | Measurement
|  |  |  | B038H | Pipe diameter
|  |  |  | B038H | Modified pipe diameter
|  |  |  | 21B5H | Unseal the electronic seal password (the high two bytes of the issued password).
|  |  |  | 001DH | Modify person's identity
Customer event log | 70C2H | Recent events |  | 0002H | YYMMDDhhmmss | √ | √ | √
|  |  |  | 001BH | Event Code
|  |  |  | 21B0H | Event Code Subdivision Event
Minute-by-minute water usage log cumulative record | 70C8H | Freeze data records for minutes |  | YYMMDD |  | √ | √ | √
|  |  |  | 70C6H | Minute-by-minute water usage log positive cumulative amount
|  |  |  | 70C7H | Minute water usage log reverse cumulative amount
Freeze data records for minutes | B06BH | Freeze data records for minutes |  | YYMMDDHHMM |  | √ | √ | √
|  |  |  | B065H | Minute-by-minute freeze on positive cumulative volume (remote reporting)
|  |  |  | B056H | Minute-by-minute freeze on reverse cumulative volume (remote reporting)
|  |  |  | B067H | Instantaneous flow rate (remote reporting)
|  |  |  | B068H | Reserve 1 byte
Daily water usage log positive cumulative amount record | B033H | Daily water usage log positive cumulative amount record |  | YYMM |  | √ | √ | √
|  |  |  | B02DH | Positive cumulative amount of daily water usage log
Daily water usage log reverse cumulative record | B034H | Daily water usage log reverse cumulative record |  | YYMM |  | √ | √ | √
|  |  |  | B02EH | Daily water usage log reverse cumulative amount
Monthly water usage log positive cumulative amount record | B035H | Monthly water usage log positive cumulative amount record |  | YY |  | √ | √ | √
|  |  |  | B02FH | Monthly water usage log positive cumulative amount record
Monthly water usage log reverse cumulative record | B036H | Monthly water usage log reverse cumulative record |  | YY |  | √ | √ | √
|  |  |  | B030H | Monthly water usage log reverse cumulative record

## Dataset ID
Data terminal
ID | name | Support function codes | Read and write attributes | describe | length | Sub-item ID (Do not upload) | Sub-item length | Sub-item description | Production mode | Customer Model
[9 empty] | local | local | remote
[9 empty] | Business data is not encrypted | Business data encryption | Business data encryption
21B8H | Communication end frame | Downward | W | Total length 20. Sub-items of this product are not parsed. | 20 | 0002H | 6 | clock | √ | √ | √
[6 empty] | 0027H | 4 | Remaining water volume (This field is reserved and will not be parsed by the table).
[6 empty] | 000AH | 1 | Overdraft status
[6 empty] | 0010H | 1 | Balance state
[6 empty] | 000DH | 4 | unit price
[6 empty] | 000FH | 4 | Remaining amount
B064H | Terminal communication identification code | Up | R | 1. The first frame contains the following: real-time data (70F2H) + daily frozen data (B05CH) + alarm status word (B070H) + 10 most recent events (70EEH). 2. If the frozen data needs to be uploaded in packets, each packet must contain B064H. | 6 | B064H | 6 | Terminal communication identification code |  |  | √
70F2H | Terminal real-time data reporting | Up | R |  | 49 | 0002H | 6 | System clock, BCD code, YYMMDDHHMMSS |  |  | √
[6 empty] | 0004H | 2 | Main voltage, unsigned integer (HEX), unit 0.001V
[6 empty] | 0006H | 2 | Backup power voltage, unsigned integer (HEX), unit 0.001V
[6 empty] | 210BH | 1 | CSQ-GPRS
[6 empty] | B079H | 5 | Positive cumulative amount, unit 10L, unsigned integer. Current positive cumulative amount of water meter.
[6 empty] | B07AH | 5 | Reverse cumulative amount, unit 10L, unsigned integer. Current forward cumulative amount of water meter.
[6 empty] | B067H | 3 | Instantaneous flow rate, unit: 10L/h
[6 empty] | 0027H | 4 | Remaining quantity, in liters (L)
[6 empty] | 70B5H | 4 | Total pump operating time, in hours, unsigned integer.
[6 empty] | 70B9H | 5 | Remaining allowance for this period, in units of 10L, unsigned integer.
[6 empty] | 70BAH | 5 | Additional usage amount for this period, in units of 10L, unsigned integer.
[6 empty] | 70BBH | 3 | The maximum daily flow rate over the past 24 hours, with the maximum daily flow rate in units of 10L/h. For example, 16 represents 160L/h.
[6 empty] | 0039H | 4 | Average daily flow rate, in L/h. For example, 1600 represents 16.00 L/h. 0x7FFFFFFF indicates an anomaly.
B05CH | Daily frozen data reporting | Up | R |  | 10 | B077H | 5 | Daily frozen positive cumulative amount, unit 10L, unsigned integer. Daily cumulative positive flow rate of water meter. |  |  | √
[6 empty] | B078H | 5 | Daily frozen reverse cumulative amount, unit 10L, unsigned integer. Daily cumulative reverse flow of water meter.
B070H | Alarm status words | Up | R |  | 3 | B070H | 3 | Alarm status word. Reported only when status changes. |  |  | √
70EEH | Recent events (10 items) | Up | R |  | 90 | 70C2H | 9 | Recent events (10 items) |  |  | √
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9
[6 empty] | 70C2H | 9

## Set parameters for the whole t
Original Goldcard Industrial Electromagnetic Water Meter:
Iranian electromagnetic water meters:
| Settings/Read: | Read-only:
| Table No. | Battery main voltage
| clock | Battery backup voltage
| Work mode | Current status word
| One-click cleanup command for production | Main control software version
| Valve/pump action configuration when the event occurs | Measurement software version
Local communication key
Remote communication key
Clear positive and negative accumulations
STS Key
Set production date
Daylight Saving Time Enabled
Daylight Saving Time Configuration
Customer communication IP address and port number
Customer Communication APN
Water meter type
Pipe diameter
Range ratio